Chapter 3

Intravascular Brachytherapy

24.1. Introduction

Coronary artery disease is most commonly treated using bypass surgery or percutaneous transluminal coronary angioplasty (PTCA).

A major problem with PTCA, however, is the high incidence of restenosis and recurrence of artery blockage at the site of treatment.

Restenosis is usually arbitrarily defined as a narrowing of the lumen by 50% or greater in diameter compared to the adjacent normal-appearing segments.

Although the occurrence of restenosis is significantly reduced by the implantation of coronary stents, the restenosis rate following balloon angioplasty in randomized trials is 30% to 40% (1).

Most restenosis after angioplasty or stenting is caused by thrombosis or blood clotting at the PTCA site, which can be prevented partially by using anticlotting drugs.

However, another process, which begins within days after angioplasty, is the neointimal growth of tissues prompted by the wound-healing process following tissue injury by angioplasty.

This component of restenosis cannot be prevented by anticoagulants or stents.

Intraluminal irradiation of coronary and peripheral arteries together with balloon angioplasty and/or stent implantation significantly lowers the rate of neointimal formation, thereby reducing the rate of restenosis to well below 10% (2).

Radiation kills cells and inhibits the growth of neointimal tissues in a manner similar to its effect on benign diseases such as keloids and heterotopic bone formation.

24.2. Treatment Volume

A. Arterial Anatomy

The arteries carry blood from the heart to various parts of the body. The main artery, the aorta, is the largest blood vessel (2–3 cm in diameter) and carries blood from the left ventricle of the heart to the branching arteries and capillaries in all the body organs and tissues, including the heart muscle. The coronary arteries are blood vessels lying on the outer surface of the heart and the peripheral arteries supply blood to other organs and tissues. The luminal diameter of arteries ranges from 3 to 5 mm initially and tapers slowly through their path length. Within this range, the peripheral arteries tend to be of larger diameter than the coronary arteries. The minimum normal artery diameter required for angioplasty and stenting is approximately 3 mm.

The inside of the arteries is lined with a layer of cells called endothelium. Next to the endothelium is the connective tissue layer, the intima, followed by layers of elastic membrane, smooth muscle cells, and elastic tissues. The outermost layer of the arteries is called adventitia, made up chiefly of collagenous fiber.

B. Angioplasty and Restenosis

The arteries can be partially blocked due to atherosclerosis or plaque formation.

Reduction of their lumen diameter compromises the flow of blood and the delivery of oxygen to the body tissues. As an alternative to a major surgical procedure such as bypass surgery, balloon angioplasty is used to dilate the lumen diameter.

This stretching action often ruptures the internal elastic lamina of the wall and causes fissures in the medial layers.

The acute risk of the angioplasty procedure is thrombosis that can be controlled by drugs, as mentioned previously.

The more protracted risk, however, is that of restenosis by neointimal hyperplasia. This process involves growth of new tissues in the cracks and crevices of the arterial wall caused by angioplastic injury. Although implantation of stent angioplasty reduces the overall rate of restenosis by approximately 50%, it does not prevent neointimal growth and may, in fact, stimulate the process.

C. Target Volume

Target volume for intravascular brachytherapy (IVBT) is confined to the region of angioplasty. Typically, it is 2 to 5 cm in length of artery and 0.5 to 2 mm in thickness of arterial wall.

Occasionally, these dimensions may be exceeded depending on the location and extent of the disease.

With 3 to 5 mm of luminal diameter, the radial range of treatment may extend as far as to about 5 mm from the center of the artery.

Because of the severity of inverse square falloff of radiation at short distances, transluminal irradiation with intravascular brachytherapy produces highly conformal dose distribution, delivering a high dose to the arterial wall while sparing surrounding normal vessels or myocardium.

Again, because of the predominance of the inverse square law effect, penetrating power of radiation, depending on energy and modality, is not critically important, except with regard to dose rate or duration of implant and radiation protection of personnel involved with the procedure.

β-Particle sources, in general, give higher dose rates and provide greater radiation protection compared to the γ-ray sources.

The depth of dose prescription for intracoronary irradiation is recommended by the American Association of Physicists in Medicine (AAPM) (6) to be 2 mm from the center of the source and for the peripheral arteries 2 mm beyond the average lumen radius.

For each case, dose distribution in at least three planes perpendicular to the catheter and along its length should be determined. In addition, average, maximum, and minimum doses within the target volume should be reported (6).

24.3. Irradiation Techniques

Intravascular brachytherapy techniques may be classified into two categories:

- temporary implants (sealed sources or liquid-filled balloons) and

- permanent implants (radioactive stents)

Each method has its advantages and limitations, but the catheter-based sealed source is the most commonly used method of treatment. It is the preferred method because of its better control of dose delivery. A variety of β- and γ-ray sources have been used for endovascular therapy, although the choice of one modality over the other is yet not clearly established. The pros and cons of a few sources and devices are discussed below.

A. Radiation Sources

Typical dosimetric requirements of a temporary intravascular implant are to

(a) deliver a target dose of 15 to 20 Gy to a 2- to 3-cm length of the arterial wall involved at a radial distance of about 2 mm from the source center

(b) minimize the dose to tissues outside the region of angioplasty, and

(c) take as little time as possible for completion of the procedure, that is, provide target dose rates on the order of 5 Gy/min or greater.

These requirements suggest the suitability of high-energy β sources such as strontium-90, yttrium-90, and phosphorus-32 or high-activity γ sources such as iridium-192.

The latter could be a high-dose-rate (HDR) afterloading unit with the source dimensions small enough to allow intravascular brachytherapy.

The β sources have several advantages over γ sources:

- higher specific activity, higher dose rate,

- longer half-life, and

- greater radiation safety for the patient as well as personnel.

The major disadvantage of β sources, however, is the extremely rapid radial dose falloff within the target region;

γ sources such as 192Ir provide relatively more uniform target dose, governed primarily by the inverse square law falloff with distance, but require high activity to yield a reasonably high dose rate (≥5 Gy/min). Consequently, radiation protection problems with such sources become more significant. Although the HDR afterloaders using γ sources could provide sufficiently high dose rate, they would require expensive shielding of the catheterization laboratories.

Table 24.1 contains a list of possible isotopes that have been or could be used for intravascular brachytherapy. The last column shows the activities required to obtain a dose rate of 5 Gy/min to a 2-cm length of a vessel at a radial distance of 2 mm from the source center.

It is seen that the γ sources, because of lower specific activity, require much higher activities than the β sources for a catheter-based intravascular procedure. On the other hand, a permanent radioactive stent using 48V requires only 1 µCi to produce the same dose rate.

Although dose rate per unit-activity favors β emitters, radial dose distribution is better for the γ sources, if it is assumed that the dose uniformity across the target volume is radiobiologically beneficial. This assumption has not been clinically validated but it seems logical, based on experience in  conventional radiation therapy.

Figure 24.1 compares radial dose distribution as a function of radial distance for some of the sources listed in Table 24.1.

Nath and Liu (7) have studied radial dose function (g) for point sources of photons and electrons using Monte Carlo simulation. Their data show that from the point of view of adequate depth of penetration for intravascular brachytherapy, photon sources above 20 keV and electron sources above 1.0 MeV are acceptable.

Table 24.1 Possible Isotopes for Intraluminal Brachytherapy

Isotope; Emission; Maximum Energy (keV); Average Energy (keV); Half-life; Activity Required. Ir-192; Γ; 612; 375; 74 d; 1.0 Ci (37 GBq). Iodine-125; X-ray; 35; 28; 60 d; 3.8 Ci (140.6 GBq). Palladium-103; X-ray; 21; 21; 19 d; 3.9 Ci (144.3 G. Phosphorus-32q); β-; 1,710; 690; 14 d; 40.0 mCi (1,480 MBq). Strontium/yttrium-90; β-; 2,270; 970; 28 y; 30.0 mCi (1,110 MBq). Tungsten/rhenium-188; β-; 2,130; 780; 69 d; 35.0 mCi (1,295 MBq). Vanadium-48; β+; 690; 230; 16 d; 1.0 µCi (37 Bq) stent. From Amols HI. Physics and dosimetry of intravascular brachytherapy. In: Thomadsen B, ed. Categorical Course in Brachytherapy Physics. Oak Brook, IL: Radiological Society of North America; 1997, with permission..

B. Radiation Delivery Systems

Irradiation of blood vessels to prevent restenosis following angioplasty has been carried out using external beam as well as brachytherapy.

Current trends favor catheter-based endovascular brachytherapy devices. While research continues to develop new sources and delivery techniques, a number of systems have become available commercially. Of these, the U.S. Food and Drug Administration has approved only a few for clinical use. A brief review of some of the available devices is presented below. For more detailed product information and specifications, the reader is referred to the respective company literature.

B.1. Cordis Checkmate

The Cordis CHECKMATE System1 consists of three components: (a) a nylon ribbon containing an array of 192Ir seeds, (b) a delivery catheter, and (c) a ribbon delivery device.

The iridium seeds are 3 mm in length and 0.5 mm in diameter. The interseed spacing is 1 mm and the outer diameter of the nylon ribbon containing the seeds is 1 mm. The number of seeds in a ribbon can be altered to provide source lengths of 19 to 80 mm. Each 192Ir seed has an activity of about 33 mCi, thus making it possible to keep the treatment time within 15 to 25 minutes. The ribbon delivery device is mounted on a cart (Fig. 24.2) and uses a hand-crank mechanism to advance the ribbon into a closed-end delivery catheter.

Although the 192Ir source provides better dose homogeneity because of the lower depth dose gradient (ratio of surface to adventitial dose), its higher activity and γ-ray energy raise radiation protection concerns for personnel.

Longer irradiation times and noncentering of the delivery catheter are other disadvantages of this device.

| View Figure | Figure 24.1. Radial dose falloff with distance for a number of sources. Doses are normalized to 1.0 at a radial distance of 2.0 mm. (From Amols HI, Zaider M, Weinberger J, et al. Dosimetric considerations for catheter-based beta and gamma emitters in the therapy of neointimal hyperplasia in human coronary arteries. Int J Radiat Oncol Biol Phys. 1996;36:913–921, with permission.) |
| --- | --- |

| View Figure | Figure 24.2. The Cordis CHECKMATE system showing delivery device on a cart. (From Ali NM, Kaluza GL, Raizner AE. Catheter-based endovascular radiation therapy devices. Vasc Radiother Monitor. 2000;2:72–81, with permission.) |
| --- | --- |

B.2. Guidant GALILEO

The Guidant GALILEO System2 uses a β source, 32P, for intravascular brachytherapy. The 32P source is hermetically sealed in the distal 27-mm tip of a flexible 0.018-inch nitino (nickel-titanium alloy) wire. A spiral centering balloon catheter (Fig. 24.3), which centers the source wire, is flexible to navigate through  the arteries and also has perfusion capabilities to limit myocardial ischemia during the procedure. The source delivery unit is an automatic afterloader device that can be controlled from a remote location to advance and retract the source wire. A treatment-planning system is provided to calculate dwell times required to deliver the prescribed dose.

| View Figure | Figure 24.3. Guidant GALILEO system: A: source wire, B: source delivery unit, and C: spiral centering catheter. (From Ali NM, Kaluza GL, Raizner AE. Catheter-based endovascular radiation therapy devices. Vasc Radiother Monitor. 2000;2:72–81, with permission.) |
| --- | --- |
| View Figure | Figure 24.4. Novoste Beta-Cath system with source train and transfer device. (Courtesy of Novoste: www.novoste.com.) |

The major advantages of the system are the centering capability of the delivery catheter, automation of the afterloader, and availability of the treatment-planning system. The disadvantage is the rapid dose falloff radially, which is a characteristic of all the β sources.

B.3. Novoste Beta-Cath

The Novoste Beta-Cath System3 uses β sources of 90Sr/90Y isotope. It is a manual afterloader device with a catheter-based delivery system. The system consists of two main components: (a) a transfer device for housing and hydraulic delivery of a radiation source train and (b) a delivery catheter to transport the source train.

The delivery catheter (Fig. 24.4) has three lumens. The first lumen is used for travel of the guidewire. The second lumen is for transport of the source train consisting of sealed cylindrical seeds of 90Sr/90Y. The source train is delivered to the distal end of the delivery catheter by applying manual hydraulic pressure through a syringe, which contains sterile water and is attached to the transfer device. The third lumen, which is also attached to the transfer device, is designed to provide an opposite hydraulic pressure for returning the source train back into the storage position of the transfer device.

Major advantages of the Novoste system is its use of the 90Sr/90Y source, which is one of the highest-energy β emitters with a long half-life (28 years). Other advantages include high dose rate (treatment time approximately 5 minutes), patient and personnel safety, and simplicity of the handheld transfer device. The major disadvantage is the lack of a catheter-centering device, which could result in extreme dosimetric hot and cold spots within the target volume.

B.4. β-Emitting Liquid-filled Balloon

An alternative to catheter-based wires and seeds is to inflate the balloon dilation catheter with β-emitting radioactive liquid. The advantages of a liquid-filled balloon are inherent source centering and dose uniformity to the vessel wall. Several β-emitting isotopes such as 32P, 90Y, and 188Re, which can be obtained in a liquid radiopharmaceutical preparation, may be used in this technique.

The major disadvantages of liquid-filled balloons include (a) higher ratio of surface/adventitial dose compared to the catheter based γ-source systems and (b) the possibility, although remote, of balloon rupture and consequently leakage of radioisotope within the patient. Of the radioisotopes mentioned previously, the rhenium-188 formulation is preferable because of the reduced radiation dose to organs such as colon and thyroid due to its rapid renal elimination in the event of balloon rupture. 188Re also has a favorable maximum β energy (2.13 MeV).

One of the liquid-filled balloon devices was developed at the Columbia University/Oak Ridge National Laboratory (8). The system uses a liquid preparation of 188Re (188Re-MAG3), which is obtained at high specific activities from a tungsten (188W) generator and delivered into a perfusion angiography balloon. A commercial system (RADIANT4) also uses a 188Re-filled balloon and is similar to conventional balloon PTCA (Fig. 24.5). It may be used before or after stent placement.

| View Figure | Figure 24.5. A: Schematic of RADIANT radiation delivery system. B: PAS RADIANT isolated β-source balloon. Top: Inflated balloon. Bottom: WRAP balloon protective sheath. (From Eigler N, Whiting JS, Makkar R, et al. Isolated liquid beta source balloon radiation delivery system (RADIANT). In: Waksman R, Serruys PW, eds. Handbook of Vascular Brachytherapy. London: Martin Dunitz Ltd.; 1998:107–110, with permission.) |
| --- | --- |

B.5. Radioactive Stents

Because permanent stents are frequently used in conjunction with balloon angioplasty, incorporation of radioactivity into the stent has been suggested to make it more effective in preventing restenosis (9,10,11,12,13). 32P, 90Y, and 48V are some of the suitable β-emitting isotopes for impregnation into the stent. The stent is rendered radioactive by activation in a cyclotron or by ion implantation with the radioisotope.

The advantage of a radioactive stent is primarily the combining of two procedures, stenting and irradiation, into one. Proximity of the radioactive source with the vessel walls may be another advantage, although the gridded structure of the stent gives rise to greater dose inhomogeneity at the vessel surface than with a liquid-filled balloon or a catheter-based source. Figure 24.6 shows peaks and valleys of dose distribution for a 1.0-µCi 32P Palmaz-Schatz stent.5

The implantation technique of a radioactive stent is the same as that required for a nonradioactive stent. Because of the very low activity (e.g., 0.5–5.0 µCi) of the β radioisotope, the radioactive stent procedure is the safest from the radiation protection point of view. However, the dosimetry is much more complicated than that of the catheter-based systems, as will be discussed later.

| View Figure | Figure 24.6. Calculated dose distribution for a 1.0-µCi 32P Palmaz-Schatz stent, 15 mm in length. Dose to tissue is that given in 14.3 days after implant. (From Janicki C, Duggan DM, Coffey CW, et al. Radiation dose from a phosphorus-32 impregnated wire mesh vascular stent. Med Phys. 1997;24:437–445, with permission.) |
| --- | --- |

P.496

24.4. Dosimetry

A. Dose Calculation Formalisms

A.1. Catheter-based γ Emitters

The AAPM TG-43 formalism (13) is generally applicable to the problem of dose calculation for catheter-based intravascular systems. For a system using photon-emitting sources, the dose D at a point (r, θ) is given by:

where Sk is the air kerma strength, ʌ is the dose rate constant, G is the geometry factor, g is the radial dose function, F is the anisotropy factor, and (ro, θo) are the polar coordinates of the reference point.

As discussed in Chapter 15, the reference distance ro in conventional brachytherapy is 1 cm. For intravascular brachytherapy, the AAPM (6) recommends ro = 2 mm. Equation (24.1) is therefore modified to incorporate this reference point specifically:

where ʌro is the dose rate constant at a reference distance of ro and gro is the radial dose function normalized to the reference radial distance ro.

Depending on the source dimensions and the location of the point (r, θ), the source may be considered as a line source or a point source. In the case of the point source, Equation 24.2 can be written as an approximation:

where Φan is the average anisotropy factor.

For a uniformly distributed line source, the geometry factor is given by:

where L is the active length of the source.

A.2. Catheter-based β Emitters

Because the quantity air kerma strength does not apply to β-emitting sources, the AAPM TG-60 report recommends the following equation for the calculation of dose at a point (r, θ) for the β sources:

where D(ro, θo) is the dose rate in water at the reference point (ro, θo). This quantity, D(ro, θo), may be determined by calibration of the β source at the reference point of ro = 2 mm and θo = π/2.

A.3. Radioactive Liquid-filled Balloons

Amols et al. (14) have calculated dose at a point P(x′, y′, z′) from a radioactive liquid-filled balloon (Fig. 24.7) by numeric integration of a Monte Carlo–generated β-dose kernel over the volume of the balloon:

| View Figure | Figure 24.7. Geometry of dose calculation at a point P(x′, y′, z′) from a balloon catheter with outer diameter Ro, inner diameter Ri, and axial length L, uniformly filled with β-emitting radioisotope. (From Amols HI, Reinstein LE, Weinberger J. Dosimetry of a radioactive coronary balloon dilution catheter for treatment of neointimal hyperplasia. Med Phys. 1996;23:1783–1788, with permission.) |
| --- | --- |

P.497

where D(P) is the dose rate (grays/sec) at point P, k(r) = dose kernel (grays per decay), r′(cm) = [(x′ - x)2 + (y′ - rsin θ)2 + (z′ - rcos θ)2]½, A/V = activity per unit volume (becquerels/cm3), and dv = r.dr.dθ.dx. The integration limits are –L/2 < x < L/2, Ri < r < Ro and 0 < θ < 2π. As seen in Figure 24.7, the central channel of the catheter contains a lumen of radius Ri to allow passage of the guidewire.

Using a 3-mm-diameter and 20-mm-long balloon filled with 90Y-chloride solution, Amols et al. (14) verified the calculated dose distribution by measurement with GafChromic film. The agreement between the measured and calculated radial dose distribution was ±6% at distances of 2.5 to 5.0 mm from the center of the catheter (or 1.0–3.5 mm from the surface of the balloon). They used the kernel for 90Y, which had been calculated by Simpkin and Mackie (15).

The dose kernel is defined (15) as “the expectation value of the spatial distribution of the energy deposited in target volumes centered about a point radionuclide source per unit mass of the target volume per decay of the point source.” Dose kernels for a number of β emitters, 32P, 67Cu, 90Y, 105Rh, 131I, 153Sm, and 188Re, have been calculated using the EGS4 Monte Carlo code (15). Some of these have potential use in intravascular brachytherapy with liquid-filled balloons.

A.4. Radioactive Stents

Dosimetric and radiobiologic characteristics of radioactive stents are complicated because of the gridded structure of the stent and the differences in dose rates delivered by stent versus an acute catheter-based irradiation. Dosimetric evaluation of clinical studies in animals by the AAPM task group (6) revealed “a confusing variety of different dose specifications in radioactive stent implantation.” The analysis of a group of cases also showed that the dose delivered per unit-activity for the same radionuclide varied by almost a factor of four. Thus, the user must exercise caution in using stent activities and dosages specified in the clinical studies reported in the literature. The AAPM recommends (6) that “for each type of radioactive stent the three-dimensional dose distributions around stents of various lengths, diameters, and activities should be carefully determined by benchmark dosimetry studies before clinical implementation.”

The dosimetric calculations of radioactive stents have been investigated by a number of investigators (16,17,18,19). A stent is modeled theoretically as a cylindrical shell of water, immersed in water, with the radioisotope (a β emitter) uniformly distributed throughout the shell. The source is divided into infinitesimal elements of volume, each representing a point source. The dose distribution of a point source is represented by the dose-point-kernel function K(, ′), which is defined as the dose at point produced by a point source of unit activity located at ′. The dose D(, t) at a point outside the stent, accumulated over a time t, is calculated by convolution:.

where a(, t) is the activity volume density. The dose point kernel is derived from Monte Carlo simulation of electron transport in water (15)..

For mathematical details of the above or similar algorithms, the reader is referred to Prestwich et al. (16) and Janicki et al. (18).

B. Measurement of Dose Distribution

Dose distribution around intravascular brachytherapy sources is best measured by film dosimetry. Because of the high dose rate and steep dose gradients near the source, the film must have very thin emulsion, slow speed, and high resolution. Radiochromic films meet these requirements and are the detectors of choice for measuring dose distribution around brachytherapy sources in a contact geometry (see Chapter 8).

GafChromic film is tissue equivalent and has a linear sensitometric (optical density vs. dose) response and a large dynamic range (e.g., several grays to obtain an optical density between 0.5 and 2.5). The film does not require processing after irradiation.

In dosimetric measurements, the film is sandwiched in a water-equivalent plastic phantom or wrapped around a cylindrical phantom, just like any film dosimetry arrangement. The exposed film is analyzed by a high-resolution scanning densitometer (spatial resolution on the order of 0.1 mm). Optical density is converted to dose based on the predetermined sensitometric curve.

C. Calibration

C.1. Catheter-based γ Sources

The strength of the γ-emitter sources is determined by measurement of exposure rate in free space at 1 m using a free-air ionization chamber or a well-type ionization chamber, which bears calibration traceable to the National Institute of Standards and Technology (NIST) for a source of the same type. The air kerma strength, Sk, at distance l is given by:

P.498

where l is the exposure rate at distance l and /e is the average energy absorbed per unit charge of ionization in air, assuming energy loss due to bremsstrahlung in air to be negligible (see section 15.2A.5)..

The calibration procedure of γ-emitting seeds for conventional brachytherapy has been discussed in Chapters 15 and 23. In the case of intravascular sources, the AAPM (6) recommends that the dose rate in water at a distance of 2 mm from the center of the source must be specified. Equation 24.2 may be used to calculate dose rate at the reference point of (2 mm, π/2). However, at this point, the data for ʌro and gro are not available. Amols et al. (20) have used extrapolated values of the line source functions given by the AAPM TG-43 (13) to calculate dose distribution at distances less than 1 cm. For 192Ir source (0.5 mm in diameter, 3 mm in length), they have compared calculated distributions with those measured with GafChromic film and shown an agreement of ±7% at distances of 1.5 to 5.0 mm. Thus, until TG-43 functions are available at distances less than 0.5 cm, the user may use the calculative or experimental methodology of Amols et al. to determine the dose rate at the reference point of (2 mm, π/2).

C.2. Catheter-based β Sources

The strength of β sources for intravascular brachytherapy is specified in terms of dose rate in water at the reference point (ro, θo). Accordingly, the reference point for calibration of these sources should be ro = 2 mm and θo = π/2. Soares et al. (21) have described a method of calibrating β sources using an extrapolation ionization chamber that has a 1-mm-diameter collecting electrode (22). The absorbed dose rate at a depth close to 2 mm averaged over a 1-mm-diameter area is determined from measurements in a A150 plastic phantom. The absorbed dose to A150 plastic is converted to absorbed dose to water using an appropriate scaling factor and density corrections (21).

C.3. Radioactive Stents

Calibration of a radioactive stent is complicated by the fact that stents are shipped under sterile conditions in acrylic cylinders of wall thickness sufficient to absorb all the β particles. Activity measurements therefore need to be made without disturbing the sterile environment. Although the β particles are completely absorbed in the acrylic shield, bremsstrahlung x-rays escape. The intensity of these x-rays is proportional to the numbers of β particles emitted, thus making it possible to infer activity of the radioactive stent by external measurement of bremsstrahlung. This method of assaying stent activity has been discussed by Larson and Mohrbacher (23) and Coffey and Duggan (24).

Relative calibration procedures based on bremsstrahlung measurements are available using a well-type ionization chamber, an NaI(Tr) scintillation counter, or a liquid scintillation counter. However, standards are needed to first calibrate the instrument. The NIST offers a reference 32P standard for activity analysis and hopefully more will be forthcoming for users to calibrate their instruments through transfer techniques. Relative calibrations require reference standards of the same radionuclide and construction. Large errors can result if one compares sources of different geometries for relative calibrations.

24.5. Quality Assurance

A comprehensive quality assurance (QA) program should ensure accurate and safe delivery of intravascular brachytherapy. In addition, it should be designed to satisfy the relevant regulations of the Nuclear Regulatory Commission (NRC) or state if it is designated as an Agreement State. Although several QA reports have been published on the use of conventional brachytherapy (25,26), the report most pertinent to intravascular brachytherapy is the AAPM TG-60 (6). Recommendations of this report are summarized below:

- Document radiation source properties.

- Develop protocols for receipt of sources, acceptance testing, and commissioning procedures.

- Develop guidelines for storage, access, and inventory of sources.

- Check physical integrity of sealed source; perform leak testing and other essential pretreatment QA procedures, depending on the source or device (e.g., remote afterloader).

- Verify source activity using in-house equipment such as dose calibrator, re-entrant chamber, etc.

- Develop methods to properly sterilize sources.

- Develop protocol for safe transportation of radioactive material.

- Develop protocol for safe disposal of sources after use.

- Ensure availability of ancillary equipment at the time of the procedure.

- Develop emergency procedures.

- Develop roles and responsibilities of each individual involved with the procedure.

- Develop a prescription form (written directive).

- Develop dose calculation check/double-check procedures.

- Verify correct choice of sources and treatment parameters for the patient in question.

- Monitor radiation levels around patient and surrounding areas during treatment.

- Perform posttreatment radiation survey of the patient and surrounding areas.

- Provide radiation safety instructions to patients with permanent implants (e.g., radioactive stent) at the time of patient discharge.

- Develop education and training program for personnel.

24.6. Recent Updates

A. Dosimetry

The AAPM Task Group No. 60 on IVBT was published in 1999. This report was updated in 2007 to Task Group No. 149 (27). The new recommendations provide a consensus report on dose calculation formalism and dosimetry parameters for commonly used IVBT systems. The reader is referred to this document for implementing these updates.

B. Drug-eluting Stents

Implantation of drug-eluting stents (DESs) has shown promise to significantly reduce the incidence of restenosis (28,29). The eluting drug has anti-proliferative properties that help in the prevention of neointimal hyperplasia. However, restenosis can still occur with DESs, although at a reduced rate compared to the bare-metal stents. Torguson et al. (30) have shown that the use of intravascular brachytherapy (IVBT) after implantation of DESs can further decrease the likelihood of restenosis. Thus, IVBT has an important role in the prevention of restenosis with or without the DES.

Key Points

- IVBT significantly lowers the rate of neointimal formation, thereby reducing the risk of restenosis.

- The target volume for IVBT is typically 2 to 5 cm in length of artery and 0.5 to 2 mm of arterial wall.

- The AAPM-recommended depth of dose prescription is 2 mm from the center of the source for intracoronary irradiation. For the peripheral arteries, it is 2 mm beyond the average lumen radius.

- High-dose-rate (e.g., 5 Gy/min or greater) β sources are preferable for IVBT over g sources.

- Catheter-based radiation delivery systems are preferable over β-emitting liquid-filled balloons or radioactive stents.

- The AAPM TG-43 formalism (with the TG-149 update) is suited for the catheter-based IVBT dose calculations. Different algorithms are required for liquid-filled balloons or radioactive stents.

- The AAPM protocol recommends the calibration of IVBT sources at a distance of 2 mm from the center of the source.

- An IVBT QA program must be instituted and implemented in accordance with the national protocol guidelines (e.g., AAPM TG-60) and the mandatory requirements of the NRC license.

- Current trends favor DESs. However, IVBT still has a role in lowering the incidence of restenosis, with or without DESs.

References

1. Fisherman DL, Leon MB, Bain DS. A randomized comparison of stent placement and balloon angioplasty in the treatment of coronary artery disease. N Engl J Med. 1994;331:496–501.

[CrossRef]

2. Waksman R. Radiation for prevention of restenosis: where are we? Int J Radiat Oncol Biol Phys. 1996;36:959–961.

P.500

3. Hall EJ, Miller RC, Brenner DJ. The basic radiobiology of intravascular irradiation. In: Waksman R, ed. Vascular Brachytherapy. Armonk, NY: Futura; 1999.

4. Fajardo LF, Berthrong M. Vascular lesions following radiation. Pathol Annu. 1988;23:297–330.

[Medline Link]

5. Reinhold HS, Fajardo LF, Hopewell JW. The vascular system. In: Altman KI, Lett JT, eds. Advances in Radiation Biology, Vol 14. Relative Radiation Sensitivities of Human Organ Systems, Part II. San Diego: Academic Press; 1990:177–226.

6. Nath R, Amols H, Coffey C, et al. Intravascular brachytherapy physics: report of the AAPM Radiation Therapy Committee Task Group No. 60. Med Phys. 1999;26119–152.

7. Nath R, Liu L. On the depth of penetration of photons and electrons for intravascular brachytherapy. Med Phys. 1997;24:1358(abst.).

8. Weinburger J, Schiff PB, Trichter F, et al. Results of the Columbia Safety and Feasibility (CURE) trial of liquid isotopes for coronary vascular brachytherapy. Circulation. 1999;100[suppl 1]:1–75.

9. Fischell TA, Kharma BK, Fischell DR, et al. Low-dose beta particle emission from ‘stent’ wire results in complete, localized inhibition of smooth muscle cell proliferation. Circulation. 1994;90:2956–2963.

[Fulltext Link]

[Medline Link]

10. Laird JR, Carter AJ, Kuffs WM, et al. Inhibition of neointimal proliferation with a beta particle emitting stent. Circulation. 1996;93:529–536.

[Fulltext Link]

[Medline Link]

11. Hehrleim C, Stintz M, Kinscherf R, et al. Pure β-particle emitting stents inhibit neointimal formation in rabbits. Circulation. 1996;93:641–645.

12. Carter AJ, Laird JR. Experimental results with endovascular irradiation via a radioactive stent. Int J Radiat Oncol Biol Phys. 1996;36:797–803.

[CrossRef]

[Medline Link]

13. Nath R, Anderson L, Luxton G, et al. Dosimetry of interstitial brachytherapy sources: recommendations of the AAPM Radiation Therapy Task Group No. 43. Med Phys. 1995;22:209–234.

[CrossRef]

[Medline Link]

14. Amols HI, Reinstein LE, Weinberger J. Dosimetry of a radioactive coronary balloon dilution catheter for treatment of neointimal hyperplasia. Med Phys. 1996;23:1783–1788.

[CrossRef]

[Medline Link]

15. Simpkin DJ, Mackie TR. EGS4 Monte Carlo determination of the beta dose kernel in water. Med Phys. 1990;17:179–186.

[CrossRef]

[Medline Link]

16. Prestwich WV, Kennet TJ, Kus FW. The dose distribution produced by a 32P-coated stent. Med Phys. 1995;22:313–320.

[CrossRef]

[Medline Link]

17. Prestwich WV. Analytic representation of the dose from a 32P-coated stent. Med Phys. 1996;23:9–13.

[CrossRef]

[Medline Link]

18. Janicki C, Duggan DM, Coffey CW, et al. Radiation dose from a 32P impregnated Palmaz-Schatz wire mesh vascular stent. Med Phys. 1997;24:437–445.

[CrossRef]

[Medline Link]

19. Duggan DM, Coffey CW, Levit S. Dose distribution for a 32P impregnated coronary stent: comparison of theoretical calculations and measurements with radiochromic film. Int J Radiat Oncol Biol Phys. 1998;40:713–720.

[CrossRef]

[Medline Link]

20. Amols HI, Zaider M, Weinberger J, et al. Dosimetric considerations for catheter-based beta and gamma emitters in the therapy of neointimal hyperplasia in human coronary arteries. Int J Radiat Oncol Biol Phys. 1996;36:913–921.

[CrossRef]

[Medline Link]

21. Soares CG, Halpern DG, Wang CK. Calibration and characterization of beta-particle sources for intravascular brachytherapy. Med Phys. 1998;25:339–346.

[CrossRef]

[Medline Link]